Recommendations regarding the requirements and applications for long-term recordings in epilepsy

Demetrios Velis, Perrine Plouin, Jean Gotman, Fernando Lopes da Silva

and the  members of the ILAE DMC Subcommittee on Neurophysiology

The  ILAE Commission Diagnostic Methods (Chair John S. Duncan)

Subcommittee on Neurophysiology Position Paper

Chair: Fernando Lopes da Silva (Netherlands).

Secretary: Giorgio Battaglia (Italy).

Members: Jean Gotman (Canada), Brian Litt (USA), Bernard Renault (France), Klaus Lennertz (Germany), Demetrios Velis (Netherlands), Sylvia Kochen (Argentina), Perrine Plouin (France), Nobukazu Nakasato (Japan).

Summary:

The purpose of this paper is to update the state of knowledge with respect to Long-term monitoring (LTM) in epilepsy and to formulate recommendations regarding the application of LTM in clinical practice. LTM is an established technique in use both in a hospital setting (2, 6,12,16,19,48) and, increasingly, in an ambulatory and more recently in a community-based setting (22,46).  There has been sufficient evidence to substantiate the claim that LTM is of crucial importance in documenting electroclinical correlations both in epilepsy and in paroxysmally occurring behavioural changes often mistaken for epilepsy (21-23). Internationally recognized neurophysiological equipment standards, data acquisition and data transfer protocols and widely accepted safety standards have made widespread access to LTM facilities in epilepsy possible (30, 40, 49).  Recommendations on efficient and effective use of resources as well as regarding training and competencies for personnel involved in LTM in epilepsy have been formulated (3,5-9,16,19).  The DMC Neurophysiology Subcommittee of the ILAE recommends use of hospital-based LTM in the documentation of seizures including its application for assessing seizure type and frequency, in the evaluation of status epilepticus, in non-invasive and invasive video/EEG investigations for epilepsy surgery and for the differential diagnosis between epilepsy and paroxysmally occurring nonepileptic conditions, in children and in adults.  Ambulatory out-patient and community-based LTM may be used as a substitute for in-patient LTM in cases where the latter is not cost-effective or feasible or when activation procedures aimed at increasing seizure yield are not indicated. However, outpatient ambulatory monitoring may be less informative than is inpatient monitoring in some cases because: 1. reduction of medication to provoke seizures may not be safe as an outpatient; 2. faulty electrode contacts cannot quickly be noticed and repaired; 3. the patient may move out of video surveillance; 4. duration of ambulatory monitoring can be limited by technical constraints.

Introduction:

The purpose of this position paper is to update the state of knowledge based on existing national and international guidelines on the application of Long-Term Monitoring (LTM) in epilepsy, and to provide a selective review of the literature on emergent issues and controversies, particularly with respect to the use of provocative techniques to increase the yield of clinically relevant seizures. From this critical review a number of recommendations regarding the requirements and applications for LTM in epilepsy can be formulated. Guidelines for LTM in epilepsy have been formulated in the past (2,16,19.) The American Clinical Neurophysiology Society’s Guideline 12 on LTM for epilepsy (originally published in (6)) is currently undergoing a major revision. Recommendations may be as regards activities and competencies of registered EEG technologists involved in such investigations (9).Advances in epileptic seizure documentation over time, particularly video/EEG seizure monitoring, have led to the present ILAE’s classification of epileptic seizures (1,4). More recently the epileptic seizures have been re-codified in terms of a glossary of terms, largely derived from long-standing use of LTM (15).

Technical note:

Advances in technology have led to miniaturisation of equipment that have made multichannel EEG recording portable enough to allow high-quality (ambulatory) recording and data transmission over long distances possible.  Similarly, technological advances have allowed for highly compressed digital video acquisition and review, thus exploiting the full power of LTM in a clinical as well as an ambulatory setting. Ground rules in hospital equipment communications protocols have been constantly changing. However, with the advent of the Health Level 7 (HL7) standard some recommendations appear to have gained acceptance among manufacturers of equipment used for clinical neurophysiological purposes in general, and LTM in particular. The American Society for Testing and Materials (ASTM) has defined the E1467-94 “Standard Specification for Transferring Digital Neurophysiological Data between Independent Computer Systems.”

Standardization of data transfer protocols has been preceded by a parallel standardization in data acquisition protocols for biological signals. This has led to the emergence of two standards, which have met with varying acceptance.  These are the E 1467-92 ASTM standard (30), a very broad standard, which has not been widely used, and EDF, the European Data Format (49), more commonly used but causing some problems because it is not a complete standard.  Most manufacturers of commercially available LTM equipment routinely provide output in the EDF format in addition to their proprietary data formats for EEG. Digital video and audio information is most commonly available in MPEG level 1 or 2 compression but the synchronization between EEG and video is usually proprietary and has not been standardized.  Digital recording equipment for scalp-recorded LTM should allow for minimally 32 channel recording, including eye movement, EKG and selected polygraphy leads depending upon the clinical problem to be addressed. A certain overlap with polysomnography electrode derivations and sensors may be advantageous in distinguishing between epilepsy and sleep disorders.  The EEG system’s analog to digital converters should allow for a minimum of 12 bit precision. Data skew is to be minimized while recording the EEG signal. Montage reformatting should be made possible for on-line monitoring and for off-line reviewing purposes. Judicious digital filtering may be performed as scalp-recorded LTM signals are commonly fraught with movement and muscular artefact. Simultaneous digital video/EEG recording usually requires a synchronization process for the correlation of digital video frames to the digitally sampled EEG.

In-patient versus out-patient LTM in epilepsy:

Specific recommendations on the organisation and staffing of clinical facilities are beyond the scope of the present position paper. Nonetheless we should emphasize that an Epilepsy Monitoring Unit (EMU) should be adequately equipped and staffed such that the staff should be able to treat a situation of status epilepticus, and facilities  should be available to quickly transfer to an intensive care unit a patient who is in a convulsive state and does not respond adequately to first line therapy. Guidelines as regards basic clinical neurophysiology and requirements may be found in (3). In-patient LTM most commonly involves admission to an EMU, most often for extracranial and sometimes for intracranial LTM recordings. Practical recommendations on establishing and operating an EMU facility may be found in (21).   Practical recommendations as regards ambulatory LTM may be found in (50).

Out-patient ambulatory LTM in epilepsy is a relatively new development.  Equipment and recording practices tend to rely on the long-standing experience with community-based monitoring in sleep disorders and actigraphy (recording of a measure of gross motor activity) studies. An early evidence-based analysis exploring the benefits and limitations of out-patient EEG appeared in (24).  There is no recently published technology assessment report for epilepsy available on the combined use of out-patient EEG and digital video.  However there is some evidence from one group of investigators that temporal lobe epilepsy may lend itself for a out-patient LTM service (46).

To be cost-effective, LTM facilities should be available and tailored to the individual needs of the patient.  In scalp LTM, electrode derivations should allow for combination of both International 10-20 system positions and selected closely-spaced electrodes according to the so-called 10% system combinatorial electrode nomenclature proposed by the American EEG society (5,8), particularly those in the inferior temporal regions. Nasopharyngeal electrodes tend to cause discomfort over time to the patient undergoing LTM and should be avoided. The use of sphenoidal electrodes in presurgical evaluation cases undergoing LTM in which clinical evidence and neuroimaging suggest possible mesial temporal lobe epilepsy is still advisable (20,34,47) although this issue is a matter of contention as inferiorly placed 10% system electrodes may yield highly visible interictal and ictal onset patterns rivalling those obtained from sphenoidal electrodes (14,32).

The EMU facility should comply with the local standards of safety with respect to use of equipment connected to the electrical mains  and making patient contact.  Especially for EMU units that are equipped to record from intracranial electrodes these standards are more stringent, usually requiring compliance with patient safety standards where electrical mains equipment makes direct contact with body fluids (40).

Although life support equipment is not routinely used in the EMU, occasional life threatening emergencies such as a (convulsive) status epilepticus may occur (43,45) and should be dealt with appropriately.

The definition of “long term” in LTM depends upon the purpose for which LTM should be used. There is no generic, a priori definable upper time limit for LTM duration in epilepsy.  There are however, certain operational limits primarily determined by whether the patient is an adult or a child, the patient’s condition, the case load and logistics of an EMU, and the willingness of an institution to cover costs incurred above and beyond a fee for LTM services agreed with a third-party payer (23).

AED withdrawal increases the frequency of seizures, and generally shortens the total LTM recording time (17). In several studies the frequency of seizures increased associated with the process of drug tapering or with the period of being effectively free of medication (AED serum level confirmed), although not all types of drugs are associated with an early increase of seizures frequency. The shortest reported average time for successful, clinically relevant LTM was 4.8 days in a group of patients with mesial temporal lobe epilepsy. Other groups reached average times ranging from 5.5 to 7.6 days (17-21). The different studies were characterized by heterogeneity in patient population, aetiology and method of AED withdrawal.

Blum (13) has made an extensive statistical study that allows estimating the number of seizures that should be recorded during LTM in an epilepsy monitoring unit to determine with a given accuracy the lateralization of seizure focus. He defined “Concordance” as the fraction of seizures which arise from the preponderant site that ranges from 0.5 to 1.0. Based on 605 seizures from 57 consecutive patients with non-lesional temporo-limbic epilepsy he found that 15% of patients have “Concordance” of less than 0.7 and 80% have “Concordance” >0.9. The observation of 5 concordant seizures implies a 95% chance of the patient having a “Concordance” of at least 0.9. If one discordant seizure is recorded, then to reach the 95% confidence level of 0.9 “Concordance” requires a total of 11 concordant seizures. In this context we should add that the fact that seizures may tend to occur in clusters merits further consideration. In a study of the cluster effect on seizure lateralization during LTM in patients with bilateral foci, Haut et al (28) reported that seizures that occur at inter-seizure intervals of less than 8 hours are more likely to occur from the same side as the previous seizures, than seizures that occur at longer intervals. Therefore these authors proposed that clustered seizures should not be given the same weight as those that occur at long (>8 hours) intervals, with respect to a decision about lateralization, and consequently more than 5 seizures may sometimes be necessary for an adequate assessment of seizure lateralization. They  mention that in a few patients, the presence of bilateral foci was not demonstrated until more than 5 seizures were recorded. In any case Haut et al (29) concluded that seizures that occurred in clusters had as important a role in localization as did non-clustered seizures.

A utility evaluation study on the effects of LTM recording on decisions affecting health care delivery to patients suffering from seizures indicated that LTM performed in 131 consecutive patients (30% presurgical evaluation cases) altered the diagnosis  in the majority of cases recorded and had major impact on patient management changes in almost three quarters of the total population studied (23).

Specific considerations for the adult population:

The primary clinical application for LTM remains documentation of seizures for diagnostic purposes.  Documentation is particularly useful and often indicated or even mandatory in the definition of the type of seizure in order to insure appropriate treatment, in the differential diagnosis of epilepsy versus other paroxysmally occurring conditions including psychogenic non-epileptic events, in the evaluation and treatment of status epilepticus or unrecognized seizures in the ICU, and in the non-invasive or invasive recording of seizure events during the preoperative evaluation for a variety of epilepsy surgery procedures (15,20).  Local and national recommendations and indications for LTM in epilepsy vary considerably as do degrees of remuneration for clinically based LTM services.

The operational definition of the lower time limit for clinical LTM is the time at which the diagnostic problem has been solved, namely when sufficiently well-documented electroclinical events have been recorded in order to answer the clinical question. Recording of a single episode may be insufficient, if the question pertains to consistency of origin of epileptiform discharges, or if a patient is suspected of having more than one type of clinical event (see above for the question of lateralization). No required number of events can be specified, since it depends on the clinical question and the body of corollary clinical evidence, for example, whether mesial temporal sclerosis is seen on the MRI.

Within these broad margins one should endeavour to efficiently use this costly resource by maximizing the chance of a high yield of electroclinical events.  Time limits are particularly of importance in presurgical investigations (12, 17, 21-23).  In individual patients who keep reliable seizure occurrence calendars these may be used in advance to plan LTM to coincide with a higher-than-average chance of seizure occurrence.  Known factors that could provoke seizure occurrence in the individual patient (e.g., in reflex epilepsies) may be advantageously employed during LTM.  Sleep deprivation and other activation provocations may take place.  Regarding sleep deprivation, Malow et al (36, 37) carried out a study of 84 patients with medically refractory partial epilepsy who underwent LTM and were assigned to a group that was sleep deprived every other night or to a group that slept normally; in this study there was no significant difference between the two groups regarding seizure frequency.  Nonetheless some patients showed a considerable increase in seizure frequency after sleep deprivation. This question needs further analysis in well defined patient groups. Pharmacological profiles using intermittent AED sampling through indwelling peripheral lines may be advantageous and judicious AED withdrawal may be particularly useful in the presurgical cases, although at a certain risk for the patient in terms of a higher chance for occurrence of complications (38, 39, 43).

Quantitative analysis of the EEG is particularly useful during LTM in epilepsy.  Automatically detected seizure events have been made increasingly reliable over time and a variety of algorithms are widely used for this purpose in intracranially recorded LTM.  There has been considerable effort invested in algorithm improvements minimizing the chance of false positives and false negatives in seizure detection during scalp LTM recordings although confirmation of the events automatically detected by a qualified human operator (be that an EEGer or Registered EEG Technologist) is recommended (26,33,42,44).  The primary benefit of the use of automatic seizure event detection in LTM in epilepsy is reducing the tedious, time-consuming, and expensive fast visual review time by the qualified human operator.

Seizure events harvested should ideally be the habitual ones in each case and if doubt arises one should verify this by reviewing the seizure semiology together with the patient or with the next-of-kin, close friends, teachers or colleagues of the patient.  Pharmacological challenge with substances known to cause seizures is not recommended, given the considerable chance that their use might lead to seizure episodes other than the habitual ones. This means that agents causing seizures should not be used.  Hypnosis and other behavioural approaches including suggestion have been reported to be effective in eliciting events in patients suspected of having a psychogenic non-epileptic attack disorder (10,11,23,25,27,35), but some carry the disadvantage of deception.

Specific considerations for the paediatric population:

Paediatric LTM in epilepsy has been effectively used in the diagnosis of seizure disorders as early in life as during the neonatal period (41). Cranial circumference varies with age and does not allow the use of the 10-20 system from birth. In the neonatal period the system has to be modified with one or two temporal electrodes at birth, increasing progressively the number of electrodes to reach the 10-20 system around the age of 3 years. In very young children collodion must not be used as the babies’ skin is very fragile: special paste which can be taken off with water is available in neonatal ICU. In older children presence of parents is very useful and should be systematic: they know how to recognize the onset of their child’s seizures, and they can help in keeping the child in front of the camera. Seizures are much more frequent in children compared to adult patients, allowing shorter recordings. In many indications except for surgery, one to two days are sufficient. In childhood epilepsy, not only seizures have to be recognized and analyzed but also syndromes, and for that purpose different types of seizures may have to be recorded; moreover specific interictal EEG patterns (wakefulness and sleep) have to be looked for as they belong to the syndromic definition. Ambulatory EEG monitoring is only recommended in first approach of differential diagnosis or in children with continuous spike and wave during sleep, without any clinical seizure.

Recommendations by the DMC Subcommittee on Neurophysiology:

Although there are no published data indicating changes in global patterns of LTM use in epilepsy, a number of factors have contributed to a higher level of utilisation of this resource.  These factors include miniaturisation, standardisation and unit cost reduction of digital EEG and digital video recording equipment. National and international recommendations have been published and revised to keep pace with technical developments over time.  Thus certain patterns of efficient and effective use of this resource have appeared in reviews of the literature indicating specific and highly rewarding applications for LTM in epilepsy, both in children and in adults.

The DMC Subcommittee on Neurophysiology recommends the use of LTM in epilepsy for the following indications:

Acknowledgements:

The authors are thankful for the discussions with the members of the Commission: Diagnostic Methods  (John S Duncan, Chair)  and their suggestions as well as

those of Solomon L. Moshé.

References:

1.	Proposal for revised clinical and electroencephalographic classification of epileptic seizures. From the Commission on Classification and Terminology of the International League Against Epilepsy. Epilepsia. 1981;22:489-501

2.	American Electroencephalographic Society guidelines for long-term neurodiagnostic monitoring in epilepsy. J Clin Neurophysiol. 1985;2:419-452

3.	American Electroencephalographic Society Guidelines in EEG, 1-7 (revised 1985). J Clin Neurophysiol. 1986;3:131-168

4.	Proposal for revised classification of epilepsies and epileptic syndromes. Commission on Classification and Terminology of the International League Against Epilepsy. Epilepsia. 1989;30:389-399

5.	American Electroencephalographic Society guidelines for standard electrode position nomenclature. J Clin Neurophysiol. 1991;8:200-202

6.	Guideline twelve: guidelines for long-term monitoring for epilepsy. American Electroencephalographic Society. J Clin Neurophysiol. 1994;11:88-110

7.	Guideline four: standards of practice in clinical electroencephalography. American Electroencephalographic Society. J Clin Neurophysiol. 1994;11:14-15

8.	Guideline thirteen: guidelines for standard electrode position nomenclature. American Electroencephalographic Society. J Clin Neurophysiol. 1994;11:111-113

9.	National competency skill standards for long-term monitoring in epilepsy. Am J Electroneurodiagnostic Technol. 2005;45:61-71

10.	Barry JJ, Atzman O, Morrell MJ. Discriminating between epileptic and nonepileptic events: the utility of hypnotic seizure induction. Epilepsia. 2000;41:81-84

11.	Bazil CW, Kothari M, Luciano D et al. Provocation of nonepileptic seizures by suggestion in a general seizure population. Epilepsia. 1994;35:768-770

12.	Benbadis SR, O'Neill E, Tatum WO, Heriaud L. Outcome of prolonged video-EEG monitoring at a typical referral epilepsy center. Epilepsia. 2004;45:1150-1153

13.	Blume WT. Prevalence of bilateral seizure foci and implications for electroencephalographic telemetry monitoring and epilepsy surgery. Electroencephalogr Clin Neurosphysiol. 1994, Nov, 91(5): 329-36.

14.	Blume WT. The necessity for sphenoidal electrodes in the presurgical evaluation of temporal lobe epilepsy: con position. J Clin Neurophysiol. 2003;20:305-310

15.	Blume WT, Luders HO, Mizrahi E et al. Glossary of descriptive terminology for ictal semiology: report of the ILAE task force on classification and terminology. Epilepsia. 2001;42:1212-1218

16.	Brodie MJ, Shorvon SD, Canger R et al. Commission on European Affairs: appropriate standards of epilepsy care across Europe.ILEA. Epilepsia. 1997;38:1245-1250.

17.	Cascino GD. Video-EEG monitoring in adults. Epilepsia. 2002; 43 Suppl 3:80-93.

18.	Claassen J, Mayer SA, Kowalski RG, Emerson RG, Hirsch LJ. Detection of electrographic seizures with continuous EEG monitoring in critically ill patients. Neurology. 2004; 62:1743-8.

19.	Engel J, Jr., Burchfiel J, Ebersole J et al. Long-term monitoring for epilepsy. Report of an IFCN committee. Electroencephalogr Clin Neurophysiol. 1993;87:437-458

20.	Fernandez Torre JL, Alarcon G, Binnie CD, Polkey CE. Comparison of sphenoidal, foramen ovale and anterior temporal placements for detecting interictal epileptiform discharges in presurgical assessment for temporal lobe epilepsy. Clin Neurophysiol. 1999;110:895-904

21.	Fitzsimons M, Browne G, Kirker J, Staunton H. An international survey of long-term video/EEG services. J Clin Neurophysiol. 2000;17:59-67

22.	Flink R, Pedersen B, Guekht AB et al. Guidelines for the use of EEG methodology in the diagnosis of epilepsy. International League Against Epilepsy: commission report. Commission on European Affairs: Subcommission on European Guidelines. Acta Neurol Scand. 2002;106:1-7

23.	Ghougassian DF, d'Souza W, Cook MJ, O'Brien TJ. Evaluating the utility of in-patient video-EEG monitoring. Epilepsia. 2004;45:928-932

24.	Gilliam F, Kuzniecky R, Faught E. Ambulatory EEG monitoring. J Clin Neurophysiol. 1999;16:111-115

25	Goldstein LH, Drew C, Mellers J et al. Dissociation, hypnotizability, coping styles and health locus of control: characteristics of pseudoseizure patients. Seizure. 2000;9:314-322.

26.	Gotman J. Automatic detection of seizures and spikes. J Clin Neurophysiol. 1999;16:130-140

27.	Guberman A. Psychogenic pseudoseizures in non-epileptic patients. Can J Psychiatry. 1982;27:401-404

28.	Haut SR, Legatt AD, O'Dell C, Moshe SL, Shinnar S. Seizure lateralization during EEG monitoring in patients with bilateral foci: the cluster effect. Epilepsia. 1997 Aug;38(8):937-40

29.	Haut SR, Swick C, Freeman K, Spencer S. Seizure clustering during epilepsy monitoring. Epilepsia. 2002 Jul;43(7):711-5.

30.	Jacobs EC, Lagerlund TD, Collura TF, Burgess RC. Data interchange for clinical neurophysiology. Stud Health Technol Inform. 1993;6:195-202

31.	Jordan KG. Nonconvulsive status epilepticus in acute brain injury. J Clin Neurophysiol. 1999; 16:332-40

32.	Kanner AM, Parra J, Gil-Nagel A et al. The localizing yield of sphenoidal and anterior temporal electrodes in ictal recordings: a comparison study. Epilepsia. 2002;43:1189-1196

33.	Khan YU, Gotman J. Wavelet based automatic seizure detection in intracerebral electroencephalogram. Clin Neurophysiol. 2003;114:898-908

34.	Kissani N, Alarcon G, Dad M et al. Sensitivity of recordings at sphenoidal electrode site for detecting seizure onset: evidence from scalp, superficial and deep foramen ovale recordings. Clin Neurophysiol. 2001;112:232-240

35.	Kuyk J, Leijten F, Meinardi H et al. The diagnosis of psychogenic non-epileptic seizures: a review. Seizure. 1997;6:243-253

36.	Malow BA, Passaro E, Milling C, Minecan DN, Levy K. Sleep deprivation does not affect seizure frequency during inpatient video-EEG monitoring. Neurology. 2002 Nov 12;59(9):1371-4.

37.	Malow, BA. Sleep deprivation and epilepsy.  Epilepsy Currents 2004, 4(5): 193-195.

38.	Marciani MG, Gotman J, Andermann F, Olivier A. Patterns of seizure activation following withdrawal of anti- epileptic medication. Neurology 35:1537-1543, 1985.

39.	Marciani MG, Gotman J. Effects of drug withdrawal on location of seizure onset. Epilepsia 27:423-431, 1986.

40.	McAllister P, Jeswiet J. Medical device regulation for manufacturers. Proc Inst Mech Eng [H]. 2003;217:459-467

41.	Mizrahi EM. Pediatric electroencephalographic video monitoring. J Clin Neurophysiol. 1999;16:100-110

42.	Qu H, Gotman J. Improvement in seizure detection performance by automatic adaptation to the EEG of each patient. Electroencephalogr Clin Neurophysiol. 1993;86:79-87

43.	Rose AB, McCabe PH, Gilliam FG et al. Occurrence of seizure clusters and status epilepticus during in-patient video-EEG monitoring. Neurology. 2003;60:975-978

44.	Saab ME, Gotman J. A system to detect the onset of epileptic seizures in scalp EEG. Clin Neurophysiol. 2005;116:427-442

45.	Sanders PT, Cysyk BJ, Bare MA. Safety in long-term EEG/video monitoring. J Neurosci Nurs. 1996;28:305-313

46.	Schomer DL, Ives JR, Schachter SC. The role of ambulatory EEG in the evaluation of patients for epilepsy surgery. J Clin Neurophysiol. 1999;16:116-129

47.	Sperling MR, Guina L. The necessity for sphenoidal electrodes in the presurgical evaluation of temporal lobe epilepsy: pro position. J Clin Neurophysiol. 2003;20:299-304

48.	Thompson JL, Ebersole JS. Long-term in-patient audiovisual scalp EEG monitoring. J Clin Neurophysiol. 1999;16:91-99

49.	Varri A, Kemp B, Penzel T, Schlogl A. Standards for biomedical signal databases. IEEE Eng Med Biol Mag. 2001;20:33-37

50.	Waterhouse E. New horizons in ambulatory electroencephalography. IEEE Eng Med Biol Mag. 2003;22:74-80

# Footnotes

Other accredited communications and messaging standard, mostly North American in origin, forming an interface between clinical neurophysiological data and neuroimaging systems include the Accredited Standards Committee (ASC) X12N, Institute of Electrical and Electronic Engineers, Inc. (IEEE) P1157 Medical Data Interchange Standard (MEDIX), Digital Imaging and Communications (DICOM) developed by the American College of Radiology and the standards defined by ASTM Subcommittees E1238, E31.11, E1394, E1467 Electrical Manufacturers' Association (ACR-NEMA).

# Text boxes

- Detection, characterization, and quantification on video/EEG of ictal events, including the appropriate activation procedures to elicit them in individual patients in whom the diagnosis of an underlying epilepsy has already been made, and when the type of seizure or syndrome is not clear.

- Documentation of the electroclinical manifestations of habitual seizures including non-invasive and invasive video/EEG LTM during presurgical evaluations.

- Differential diagnosis between epileptic and nonepileptic conditions, characterized by frequently and intermittently occurring behavioural changes including psychogenic nonepileptic events and sleep disorders, particularly those involving paroxysmal movement disorders.

- Documentation of diurnal or circadian variation in occurrence of epileptiform paroxysms, in conjunction with pharmacological interventions and/or of the effect of these interventions on diurnal or circadian behavioural changes.

- Documentation of specific patterns in the occurrence of epileptiform paroxysms during sleep and/or of disruption of sleep architecture in so-called “cognitive epilepsy” cases in the pediatric population.

- Monitoring in the intensive care unit (ICU) for the effectiveness of treatment of status epilepticus and for the identification of subclinical seizures and subclinical status epilepticus, conditions that have been shown to be more frequent than usually thought in the ICU (18,31).